## Slide 1
Sistemi di localizzazione 3D
Vincenzo Ferrari
vincenzo.ferrari@unipi.it

## Slide 2: Marker-Based Tracking for Image-Guided Interventions
Ziv Yaniv
Sheikh Zayed Institute for Pediatric Surgical Innovation
Children's National Medical Center
Last updated: Sep. 17 2012

## Slide 4: GPS is not for the Operative Room

## Slide 5: Laparoscopic Navigator
EndoCAS – University of Pisa, Italy
G. Megali,V. Ferrari, C. Freschi, B. Morabito, F. Cavallo, G. Turini, E. Troia, C. Cappelli, A. Pietrabissa, O.Tonet, A. Cuschieri, P Dario F. Mosca: EndoCAS navigator platform: a common platform for computer and robotic assistance in minimally invasive surgery. The international journal of medical robotics computer assisted surgery: MRCAS 2008;4(3):242-51.

## Slide 6: Laparoscopic Navigator

## Slide 7: Motivation
Display a dynamic virtual replica of the real world on screen, facilitating understanding of current spatial relationships between objects and enabling the clinician to predict the consequences of motions.

## Slide 8: Motivation
Display a dynamic virtual replica of the real world on screen, facilitating understanding of current spatial relationships between objects and enabling the clinician to predict the consequences of motions.
Automated tool positioning (robot end effector pose) using a prediction model based on current and previous poses of other tools and anatomical structures.

## Slide 9: Motivation
Display a dynamic virtual replica of the real world on screen, facilitating understanding of current spatial relationships between objects and enabling the clinician to predict the consequences of motions.
Automated tool positioning (robot end effector pose) using a prediction model based on current and previous poses of other tools and anatomical structures.
User interaction.

## Slide 11: Spazi di lavoro e sottomoduli
WorkSpace Localizzatore
Sistema di localizzazione
Il localizzatore consente di effettuare misure 3D della sala operatoria
WorkSpace robot
Strumenti Robotici
WorkSpace sala operatoria
Paziente Reale
Strumentazione tradizionale
WorkSpace dispositivo di imaging
Immagini Pre/Intraoperatorie – “Paziente Virtuale”
Calibrazione: determinare la rototraslazione che permette di sovrapporre il workspace di dispositivi rispetto a sistemi di riferimento noti
Registrazione: determinare la trasformazione che consente di sovrapporre il “paziente virtuale” con il “paziente reale”

## Slide 12: Mechanical Systems
Forward/Direct kinematics approach.
Encoders at the joints provide the relative translation and rotation.
Horsley-Clarke Frame
Faro Arm
Da-Vinci Robot
Refresh rate and latency | ++
Concurrency | ---
Working volume | -
Obtrusiveness | --
Completeness | +
Accuracy | ++
Robustness | ++
Cost | -
Pro: Highly accurate
Con: Single object at a time
Esistono sia attivi (robot) che passivi

## Slide 14
Horsley-Clarke Frame

## Slide 15
Dime chirurgiche

## Slide 16: Surgical guide for dental implantology

## Slide 17: Surgical guide for dental implantology

## Slide 19
Ultrasonic Systems
Trilateration approach.
Compute an emitter’s distance from a set of receivers at known locations. Distance obtained using the time it takes sound to reach the receiver from the emitter (TOF).
microphones mounted on operating light:
“A Frameless Stereotaxic Operating Microscope for Neurosurgery”, E. M. Friets et al., IEEE Trans. Biomed. Eng., vol. 36(6), pp. 608-617, 1989.
Refresh rate and latency | -
Concurrency | -
Working volume | +
Obtrusiveness | +
Completeness | +
Accuracy | +
Robustness | +-
Cost | -
Pro: Unobtrusive.
Con: Low refresh rate.

## Slide 20: Ultrasonic Systems
A set of receivers (n>2) and an emitter yield a set of n quadratic equations:
m1
m2
m3
p
d1
d3
d2
In 3D when we have 3 receivers we can obtain two linear equations and a quadratic by subtracting the first equation from the other two:
Solve the two linear equations and obtain the x and y coordinates as a function of the z coordinate. Substitute back into the quadratic equation and obtain the z coordinate.
“Efficient solution and performance analysis of 3-D position estimation by trilateration”, D. E. Manolakis, IEEE Trans. Aerosp. Electron. Syst., vol. 32(4), pp. 1239-1248, 1996.

## Slide 21: Ultrasonic Systems
Three possible solutions for z:
  Imaginary solution, no intersection between the three spheres.
  One solution, single intersection point.
  Two solutions, two intersection points.
For n>3, create a set of linear equations and solve either in an exact manner (n=4) or using a least squares formulation (pseudo-inverse).
Mount 3 or more emitters on tracked object in a known spatial configuration, fire them one after the other with a minimal temporal separation that is greater then the time it takes to reach all receivers.
Once all emitters are localized estimate the transformation using analytic paired point registration.

## Slide 22: Electromagnetic Systems
Inducing current via magnetic field.
AC or pulsed DC driven magnetic fields.

## Slide 23: Electromagnetic Systems
Set of coils inside electromagnetic field generator.
Miniature coil(s) embedded in tools (5DOF or 6DOF).
Pulsed DC: sequence of static fields, pose obtained by measuring magnitude of sensor response to each field – problems with ferrous materials.
AC: repeated signal pattern on high frequency carrier wave, pose obtained by amplitudes of the received signals – problems with eddy currents and resulting magnetic fields.
There is a wireless transponder based system which is part of a radiotherapy system(Calypso Medical), but most are wired.

## Slide 24: Electromagnetic Systems
Refresh rate and latency | +
Concurrency | +
Working volume | +
Obtrusiveness | + [depends on FG choice]
Completeness | +
Accuracy | +
Robustness | +
Cost | +
Pro: Do not require line of sight.
Con: Wired, and susceptible to transient distortions.

## Slide 25: ESERCIZI
Definire una trasformazione a partire dalla conoscenza di 3 punti
Definire un sistema di riferimento note le matrici di trasformazione di 2 sensori a 5dof